## Sheet1
Audit Item | Yes | No | Comments/Action | Notes:
Precleaning |  |  |  | SDS + IFUs readily accessible to staff
Precleans the device at the point of use. |  |  |  | Standardize periods vs no periods in each statement
Utilizes precleaning solution and tools as recommended per IFU
Flushes or suctions lumens per IFU
Discards the cleaning solution and cloth after use.
Transporting
Transports the contaminated device and accessories to the device processing room as soon as possible after use.
Ensures the container or cart is labeled with a biohazard legend.
Transports immediately after pre-cleaning
Leak Testing
Performs leak testing before manual cleaning if indicated.
Ensure leak tester used is the same leak tester recommended by the IFU
If indicated, perform a dry leak test prior to device immersion
Ensures caps and covers are put on to prevent damage from immersion
Pumps the leak tester to recommended pressure per IFU
Fully immerse the device in water (no enzymatic)
Articulates distal tip as recommended per IFU
Maintains proper times for leak testing
Remove from water
Depressurize the scope
Verbalizes process for failed leak test
Manual Cleaning
Uses a freshly prepared cleaning solution and does not add additional products to the water unless recommended by the manufacturer.
If enzymatic is indicated, measures amount of enzymatic and water to achieve correct solution ratio
Completely submerges the device and accessories
Cleans exterior surfaces of the device with a soft, lint-free cloth or sponge.
Cleans all accessible channels and the end of the device with a cleaning brush of the length, width, and material recommended by the device manufacturer.
Uses a clean brush for each device cleaning.
If the endoscope has an elevator, raises and lowers the elevator throughout the manual cleaning process.
Brushes the accessible channels until no debris appears on the brush.
Removes debris before retracting the brush back through the device.
Flushes the channels of the device with the cleaning solution.
Manually actuates the valves during the cleaning process.
Flushes and rinses exterior surfaces and internal channels with water until all cleaning solution and residual debris is removed.
Dries exterior surfaces and removable parts of the device and purges all channels with air.
Reprocesses reusable parts, accessories, and cleaning implements according to the manufacturer’s instructions for use (IFU).
Disposes of single-use parts, accessories, and cleaning implements.
Inspection
Inspects and evaluates devices and accessories for:
cleanliness
missing parts
clarity of lenses
integrity of seals and gaskets
physical or chemical damage
moisture
function
If present, ensure tip of device is not damaged a wipe if indicated
Uses additional illumination and magnification for inspection, as needed.
High-level Disinfection or Sterilization
Manually cleans the device and accessories before mechanical or manual high-level disinfection or sterilization.
Mechanical methods
Checks the expiration date of the high-level disinfectant or liquid chemical sterilant before each use.
Uses a test strip or other FDA-cleared testing device specific to the disinfectant or liquid chemical sterilant and minimum effective concentration of the active ingredient for monitoring solution potency before each use.
Positions devices and accessories within the mechanical processor to ensure contact of the processing solutions with all surfaces of the device.
Connects the device to the mechanical processor correctly.
Verifies mechanical processing cycles are completed as programmed.
Manual methods
Checks the expiration date of the high-level disinfectant before each use.
Ensure test strip bottle has been validated per MIFU
Uses a test strip or other FDA-cleared testing device specific to the disinfectant and minimum effective concentration of the active ingredient for monitoring solution potency before each use.
Checks temperature of high level disinfectant to ensure it is within specified range
Flushes and fills lumens and ports with the high-level disinfectant.
Completely immerses the device in the high-level disinfectant solution for the designated time according to the device and high-level disinfectant solution manufacturer’s IFU.
Set a timer per the manufacturer's IFU
Rinse
Dons a clean set of PPE prior to touching the high level disinfected device
Rinses the device with water that meets the manufacturer’s specification or as recommended by professional organizations after disinfection.
If indicated perform multiple rinses per manufacturer's IFU
May be required for both mechanical and manual methods
Flushes lumens using 70% to 90% ethyl or isopropyl alcohol according to the device manufacturer’s IFU.
Dries exterior surfaces and removable parts of the device and purges all channels with air.
Storage
Wears clean gloves when transporting the device to and from the storage cabinet.
Device is stored in a manner that protects from contamination and growth of microorganisms (If flexible endoscope) Based on the cabinet design, stores flexible endoscopes horizontally or hangs them vertically so they do not coil or touch the floor of the cabinet.
Stores the device with all valves open and removable parts detached.
Devices stored per facility policy to communicate HLD date and clean vs dirty
Records
Processing records include
date and time
identity of endodevice and endodevice accessories
method and verification of cleaning and results of cleaning verification testing
number or identifier of the mechanical processor or sterilizer and results of process efficacy testing
identity of the persons performing the processing
lot numbers of the processing solutions
disposition of defective items or equipment
maintenance of water systems, devices and device accessories, and processing equipment
Procedural records include
date and time
identity of the patient
procedure
identity of the licensed independent practitioner performing the procedure
identity of the device and device accessories used
Training
Training is documented annually and more frequently as needed
Monthly competencies are preformed for at least 1 staff member each month and every staff member yearly
Room Set up |  |  |  | Notes
Flow of work is from dirty to clean |  |  |  | Eyewash station
Dirty and clean are clearly separated |  |  |  | Ventilation/temp/ humidity parameters
Negative air flow in soiled room
Air exchanges- 10 per hour with at least 2 being outside air?
Exhaust is vented directly outside
Note on A5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Keep specific to flexible endoscopes? Reply: Changed most of them to device except when it needs to say endoscope